Policy Statement B-218

Attachment 2

PNEUMATIC TUBE CLEANING PROCEDURE LEVEL III

PROCEDURE:

- Make sure all equipment is shut down. (System Halt) at

Main Computer.

- All necessary protective clothing shall be worn by staff.

- All tube runs/transfer units will be cleaned with a 50% bleach solution and dried.  A cleaning carrier shall be run through the system repeatedly until it returns clean.

- Stations in the spill shall be cleaned (Hopper, etc.) with a 50% bleach solution and allowed to sit for 10 minutes.  Contaminated carriers will be bio-bagged and disposed of by CSS.

- Cleaning ram shall be disinfected and cleaning foam ends will be discarded and replaced with new after each use.

- Complete Laboratory Sample Incident Report (Form 210-a) upon completion of cleaning.  Notify Security and the Service Center (159 Number) of completion.